## Slide 1
z electrode piezo 1
(smaller piezo)
z electrode piezo 1
(larger piezo)
High voltage signal
High voltage signal
Jumper for smaller piezo
Top of scanner with cover removed
Schematic of circuit board with jumpers on top. By moving the right jumper one connector to the left, the z voltage on the larger piezo is grounded.
The large scanner in the MI consists of two independent piezo crystals. By grounding the z-voltage on one of the piezos, the z-range can be reduced from ~4.5 to ~1.5 micrometer to obtain a higher z-resolution when the full height is not needed. The cover on top of the scanner must be removed by unscrewing the two "inbussen" at the top. Below is a schematic of the scanner and the jumpers. For correct height measurements, another calibration file (AFML325 small z ??) must be used.
Jumper for larger piezo