Calibrating  cantilevers using the thermal noise method

The total energy that is contained in the thermal (undriven) motion of (each mode of) a cantilever is always equal to the thermal energy. The amplitude of the motion is related to the energy via the cantilever’s spring constant. In a formula:

12kxRMS2=12kBT

with

xRMS 	Root mean square deflection of the cantilever in m

k	spring constant in N/m

kB	Boltzmann constant, 1.38*10-23 J/K

T	Temperature in K

Since we can measure the deflection using the AFM, this calibration should be very easy to do. Unfortunately, DI has waited for 10 years to embed this into their software, and only the PicoForce has it. But there is a way around this. Here’s the recipe:

Method 1, or Hutter-Bechhoefer method

- Get the Stanford Research 770 FFT analyzer of the interface physics group. It’s usually in 706, but sometimes in the meethal downstairs. It has a carrying handle, but be careful, it’s quite heavy and costs about € 9500.

- Attach the A input of the SR770 to the In0 output of the break-out box.

- Calibrate the deflection measurement using the NanoScope software. Note down the deflection sensitivity S (in nm/V). Then make sure your cantilever is far away from the surface. For optimal signal to noise on the break-out box the software should be set to contact mode.

- Make sure the spectrum analyzer is set to measure to PSD (Meas->Menu->PSD), and the units are Volts RMS (Meas->Unit->VoltRMS), and the input is A (Input ->A) (with a range somewhere between 0 and -20dB, Input->Range). You should see something like this (for a contact mode tip in liquid):

If you want, change the averaging parameters and frequency range through the Average and Frequency menus. The terms Linear and Exponential stand for single-run or continous running average.

- The quantity you are looking for is the RMS area underneath the first peak. This can be found using the ->Analyze ->Band feature on the SR770. Set the start frequency and span such that as much as possible of the first mode peak is included, without including any area from low-frequency noise peaks or higher order modes

- In the top left corner you will see the text “Band= …. V”. This is the number you need, let’s call it Z.

- If you have done all these steps, calculate the force constant as follows: k=kBT(c∗S∗Z)2

Here, c is a correction factor that accounts for the different shapes that a freely vibrating and end-loaded cantilever have. For a small laser spot at the end of the cantilever it is 1.09, at 80 percent of the cantilever length it’s 1.05. Larger laser spots (relative to cantilever size) also decrease the factor [2,3,4]

NOTE 1

What I described above is the quick and dirty way of doing this. With this method you still have problems with instrument noise and the motion in higher modes of the cantilever. A more precise method to calculate the area underneath the first mode is to save the PSD. You can do this onto a floppy disk using Store/Recall->Save data->[enter filename]->Save Ascii, or use the GPIB connection and LabView program to save. Then you have to fit a harmonic oscillator behavior to the data, using e.g. Origin. Fitfunctions are available on the AFM drive, at the same location as this file. You can then calculate the area from the fit, and have less problems with the higher modes and instrument noise.

Method 2 or Sader Plan View method.

For rectangular cantilevers you can also use the data to fit a second, independent method to determine the spring constant, the so-called Sader Plan View method [5].). This method has the advantage of not depending on the deflection sensitivity, so you do not have to push your tip hard into the surface to calibrate this.

From the previous method, follow the instructions 1), 2), and 4) (skip 3), and save the data as described in NOTE1.

A) 	Import in Origin and fit a harmonic oscillator model to the data, using drivenSHO.fdf or (for stiffer cantilevers where the noise matters too) drivenSHOwithnoise.fdf, to be found in this directory. Note down the resonance frequency and the quality factor from the fit.

B)	Make sure you know the length and width of the cantilever. Use the manufacturer’s specs (for length and width they are usually accurate), or measure them in the SEM.

C)	Open Matlab, make sure this directory is the current directory,  and type “SaderPlanview”

D)	Follow the instructions, and out rolls your spring constant

NOTE 2

The SR770 cannot measure frequencies higher than 100 kHz.. We do have a spectrum analyzer that goes to higher frequencies (e.g. for stiff TM tips in air), but it cannot be read out so easily. Contact Allard for this.

1 Hutter & Bechhoefer, Rev. Sci. Instr. 64, 1868 (1993)

2 Butt & Jaschke, Nanotechnology 6, 1 (1995)

3 Walters et. al., Rev. Sci. Instr. 67, 3583 (1996)

4 Proksch et.al. Bioph. J. 84 (2): 471A-471A Part 2 Suppl. S

5 Sader, Chon & Mulvaney, Rev. Sci. Instr 70, 3967