- USER-MANUAL

- TLC10-3

1	Safety and warnings	3

2	Warranty	3

3	EC declaration of conformity	4

4	Precautions and hazards	5

5	Installation	5

5.1	Unpacking	5

5.2	Placement and ventilation	6

5.3	Initial use	6

5.4	Bath liquids	7

6	Connecting	8

6.1	Cleaning	8

7	Introduction to the TLC series	8

7.1	General	8

7.2	Pump	8

7.3	Temperature control	9

8	Operation	9

8.1	Switching on	9

8.2	Caution with powering on/off	9

9	Operating the controller	10

9.1	Overview menu items	10

9.2	Display	11

9.3	Backside apparatus	12

9.4	Safety thermostat	12

9.5	Quick start	12

10	Menu items	13

10.1.1	Menu item "Setpoint"	13

10.1.2	Menu item "Offset"	13

10.1.3	Menu item "Max Power"	13

10.1.4	Menu item "Boost heater"	13

10.1.5	Menu item "Time const"	13

10.1.6	Menu item "Stirrer"	13

10.1.7	Menu item "Low alarm"	13

10.1.8	Menu item "High alarm"	14

10.1.9	Menu item "PID parameter"	14

10.1.10	Menu item "Backlight"	14

10.1.11	Menu item "Temp units"	14

10.1.12	Menu item "Baudrate"	14

10.1.13	Menu item "Restart"	14

10.2	PID configuration	15

10.3	Manual tuning	15

Tuning can be done automatically or manually.	15

10.4	Draining bath fluid	16

10.4.1	Using the drain tap	17

10.4.2	Fast method:  drain by hand	17

10.5	Maintenance	18

11	Trouble shooting	18

11.1	General	18

12	Specifications	19

12.1	Technical specifications overview	19

12.2	TLC10-3 performance	20

12.2.1	Stability	20

12.2.2	Cool down	20

13	RS232 Communication	21

13.1	Commands	21

13.2	RS 232 Cable	22

14	Spare parts	23

14.1	TLC10-3	23

# Safety and warnings

Make sure before installing or operating the equipment to read and understand all instructions and safety precautions listed in this manual. If there are any questions concerning the operation of the equipment or about the information given in this manual please contact your local dealer or our sales department first.

Performance of installation, operation, or maintenance other than those described in this manual may result in a hazardous situation and may void the manufactures warranty.

Never operate equipment that is not correctly installed. Unqualified personnel must not operate the equipment. Avoid damage to the equipment, or its accessories, caused by incorrect operation.

Important:

- When performing service, maintenance or moving the apparatus, always disconnect the line cord of the apparatus,

- Proper skilled and trained personnel are only allowed to operate this equipment,

- Take notice of warning labels and never remove them,

- Refer service and repairs to qualified technician,

- If a problem persists, call your supplier or Tamson Instruments B.V..

# Warranty

Tamson Instruments B.V. warrants that all their manufactured equipment is free from defects in material and workmanship, preventing the machine from normal operation. Tamson Instruments B.V. does not warranty that the equipment is fit for any other use than stated in this manual. The manufacturer can only be held responsible for the security, reliability and performance of the equipment, when operated in accordance with the operating instructions, extensions, adjustments, changes and/or if repair is performed by Tamson Instruments B.V. or authorized persons only. This warranty is limited to one year from the date of invoicing. All equipment and materials are subject to standard production tolerances and variations.

# EC declaration of conformity

Manufacturer:	Tamson Instruments bv

van 't Hoffstraat 12

2665 JL  Bleiswijk

The Netherlands

Product:	Thermostatic bath and circulator

Model:		TLC10-3

The products to which this statement relates, is manufactured and dully carried out in compliance with the provisions of directive 73/23/EEC on the approximation of the laws of the member states relating to low voltage.

The products are in conformity with the following specifications:

| Item | Reference | Description | Test result |
| --- | --- | --- | --- |
| 1 | EN55016-2-1 | Conducted emission | p |
| 2 | EN61326+A1 | Conducted discontinuous emission | p |
| 3 | EN55016-2-3 | Radiated emission | p (anechoic room) |
| 4 | EN61326+A1+A2+A3 | Radiated emission | p (anechoic room) |
| 5 | EN61000-3-2 | Harmonics | p |
| 6 | EN61326 +A1+A2+A3 | Flicker | p3 |
| 7 | EN61000-4-2 +A1+A2 | ESD | p |
| 8 | EN61000-4-3 +A1 | Radiated immunity | p |
| 9 | EN61000-4-4+A1+A2 | Electrical Fast Transients | Minimum requirements pass |
| 10 | EN61000-4-5+A1 | Surges | Minimum requirements pass |
| 10 | EN61000-4-6+A1 | Conducted immunity | p |
| 11 | EN61000-4-11 +A1 | Voltage dips and Voltage variations | p |

p 	= Pass

p3	= Pass, condition of operating during Pst measurement: Operational with heating element 1500W. Pst and Plt are not evaluated in accordance with A.5 of Annex A of EN 61000-3-3(1995) + A1(2001).

not applicable were:

Conducted discontinuous emissions (Clicks)

Radiated emission (OATS)

Power Frequency Magnetic Field*

2007,  Tamson Instruments bv, The Netherlands,

Name			:	R.C. van Hall

Function		:	Director

# Precautions and hazards

Before attempting to operate the bath read all parts of this manual carefully to insure smooth operation and avoid damage to the equipment or its accessories.

If a malfunction occurs, consult section "Trouble shooting", page  18

If problem persists, call your supplier or Tamson Instruments B.V. Never operate the equipment if not correctly installed. The equipment must be operated only by qualified personnel. Avoid damage to the equipment or its accessories through incorrect operation.

READ

CAREFULLY

# Installation

Tamson Instruments B.V. is not responsible for any consequential damage or harm caused by using this bath. Repairs on the electrical system of the bath may only be carried out by well trained and authorized persons.

## Unpacking

To avoid damage during transport all Tamson baths are carefully packed for shipment. Check the packaging for external damage and make a note on the shipping documents if any damage is found. Always retain the cartons and packing material until the bath has been tested and found in good condition. Transport companies generally will not honor a claim for damages if the respective box is not available for examination. The shipment contains at least the bath/circulator. The consignment may contain other parts, individually packed in small boxes. Please see packing list for details concerning total contents of consignment.

Before filling the bath remove any remaining packing material from its interior. The interior of the bath can be accessed by taking off the lid on the top of the bath.

## Placement and ventilation

The bath has to be placed in a well ventilated area.

Do not place the bath in a dusty environment. Dust will block the condenser unit inside the apparatus which will lead to severe mechanical damage.

20 CM

10 CM

10 CM

Air circulation has to be enabled by 10 cm of free space at both sides and 15 to 20 cm at the rear of the bath.

AIR FLOW

## Initial use

Due to transportation allow the bath a 24 hour period for stabilization. This applies also when the apparatus has been tilted or fallen.

Do not switch on power, because lubrication-oil inside the compressor system has to run into the capillary.

It will take several hours before the oil has flown back into the compressor unit. When the apparatus is immediately turned on after it has been tilted, severe damage may occur to the compressor unit due to insufficient lubrication.

VISCOSITY

< 20 cSt

FLASHPOINT

>  OPERATING

TEMPERATURE

Bath liquids

The bath must be filled with a liquid suitable for the minimum operating temperature.

It is very important to select a liquid with a viscosity of less than 20 cSt at the operating temperature and a flash point which is well above the operating temperature.

METHANOL

Methanol

As bath liquid methanol or anti freeze (mixture of water and ethylene-glycol) can be used. Do not use destilled or tap water at temperatures below 5°C/41°F

Methanol is extremely flammable and can cause fire hazard. Take all necessary precautions to reduce fire hazard. When using methanol all vapors must be removed by using appropriate air ventilation.

Methanol is toxic and can cause health risks. Use appropriate ventilation and other precautions to prevent inhaling toxic vapors. If ventilation is insufficient the risk of explosion hazards can occur!

The supplier of the bath liquid (methanol) will be able to hand over all chemical details and safety precautions related to the use of methanol. These precautions must be followed when operating the bath.

As alternative to methanol a mixture of water and anti-freeze (ethylene-glycol) can be used.

| Working temperature [C] | Mixture (volume percent) |
| --- | --- |
| 5 | 90% water, 10% ethylene-glycol |
| 0 | 80% water, 20% ethylene-glycol |
| -10 | 70% water, 30% ethylene-glycol |



WELL EARTHED

Connecting

Before plugging the cooler into the mains socket, make sure the voltage of the bath corresponds to the local voltage and frequency.

Use a mains supply that is well earthed, clean of interference and suitable for the acquired electrical load of the bath.

## Cleaning

Regularly check the apparatus and condenser unit for dust. Remove dust with a vacuum cleaner.

# Introduction to the TLC series

The TAMSON model TLC circulators are designed to perform accurate temperature control required for general laboratory. The Tamson circulator is optimized for temperature control of applications requiring a high degree of stability over a broad temperature range.

## General

The TLC apparatus consists of a combination of a cooling system and a microprocessor controlled heating element. This design ensures a high degree of accuracy and reproducibility of temperature controls.

The TAMSON baths are constructed throughout from corrosion-resistant – stainless steel and Teflon – materials. The bath is effectively insulated against heat loss by a layer of Armaflex® rubber between the inner tank and outer casing.

## Pump

A circulation pump is built-in to guarantee an uniform temperature distribution within the bath and providing the possibility to circulate through a closed external system.

Standard pump: a circulation pump is provided with a capacity of 7 liters / min and a max pressure of 2,7 m head of water [2,7mtr H2O or 270 mBar].

## Temperature control

Temperature control and setting the bath temperature is regulated using a PT-100 temperature probe Class A connected to a microprocessor module. The advanced electronic control system continually computes the energy input required for optimal temperature accuracy and stability. Temperature read-out is on a 2 x 16 caracter LCD. The actual set point is computed within the controller with an accuracy of 0.001°C.

# Operation

## Switching on

If the bath has been properly filled with fluid it can be switched with the mains switch located on the front panel. Choose a working temperature (set point).

WAIT 10 MIN

BEFORE

TURNING ON

Caution with powering on/off

Be careful and do not toggle with the on/off switch. To start the cooling compressor high currents are needed which will heat the compressor motor internally. Switching off and on the compressor several times will lead to mechanical damage.

When switching off the apparatus, wait for 10 minutes before switching the system back on again.

1x

4

1

5

3

2

6

7

Operating the controller

- LCD display

- Over temperature indicator	(Red)

- Level indicator			(Blue)

- Error				(Yellow)

- Heater indicator		(Green)

- Mains switch

- Turn-push button

Use well grounded mains. Before plugging the circulator into  mains socket, make sure the voltage of the bath corresponds to the local voltage.

The front panel layout shows the turn-push button:

Next / increase:			Turn right

Previous / decrease:		Turn left

Select:				Press

## Overview menu items

-	Setpoint

-	Offset (press: <-5.00 .. +5.00°C resolution 0.01°C)

-	Max Power (press: low, med, hi, max)

-	Boost heater (press on / off)

-	Time const (press: fast, medium slow, precise)

-	Stirrer

-	Low alarm

-	High alarm

-	PID parameter

(PID set 1, PID set 2. PID set 3, PID set 4)

Proportional band (Pb=1/P where P is proportional value)

Integral

Differential

-	Backlight

-	Temp units

-	Baudrate

-	Restart

1

2

Display

3

4

1	Temperature readout

2	Applied percentage of power

3	Operating mode

4	Indicator, alarm high, alarm low, control stable

Ad 1:	When the controller starts or is restarted, the displayed value increases to a stable readout appears after a few seconds.

Ad 2:	The controller calculates every second the amount of power which should be applied for stable control. The value is displayed with a resolution of 0.1% and ranges from 0% to 99.9%.

Ad 3:	Boost		Bath is heating to set point

Cooling	Bath is cooling down to set point

Tuning Ratio	Bath is tuning for power needed at set point, first step

Tuning SA	Bath is tuning, second step

PID SP=25.00	Bath is controlling, set point is 25.00°C (example)

Ad 4:

-		Bath control is stable

-		Alarm high, press button to reset

-		Alarm low, press button to reset

## Backside apparatus

On the top - backside of the bath following can be indicated:

2

1: Safety thermostat (over temperature protection)

2: Fan

3: RS232, female

4: Motor fuse

Outlet pump

Inlet pump

1

3

4

## Safety thermostat

When the bath temperature becomes higher than the set point temperature of the safety thermostat (Indicated “1”, see drawing), the heating electronics and controller are switched-off. The red LED indicator on the front panel will light. When the temperature of the bath has been lowered with approximately 20°K, the thermostat re-sets itself automatically, however to continue normal operation the bath has to be switched off and on again. Wait for about 10 minutes to switch the bath back on again.

## Quick start

To start operating the bath in a quick way do the following:

Fill the bath with fluid

Place the power plug,

Switch the bath on using the mains switch,

Select appropriate set point temperature,

PID settings

All measuring results have been acquired using following PID settings:

Pb	:=	25

I	:=	16

D	:=	0

Under different settings its possible to achieve even better values by trimming the PID settings.

# Menu items

Use the turn-pus button to select a menu item and select the item by pressing the button. After pressing a     sign appears next to the value indicating the value can now be changed. Pressing the button again activates the value immediately. When the value is altered but the button is not pressed the value will be accepted and stored after 5 seconds automatically. The menu returns back to normal operating mode.

### Menu item "Setpoint"

Select the set point temperature. Resolution is 0.01°C +/- the system accuracy.

### Menu item "Offset"

The temperature displayed can be increased or decreased with an offset ranging from +5.00 down to -5.00 °C in steps of 0.01°C. This way the temperature reading on the display can be synchronised with an independant separate thermometer.

### Menu item "Max Power"

(press: low 25, med, hi, max)

Limits the applied power by a maximum value:

Low			Maximum of 25% applied

Medium			Maximum of 50% applied

High			Maximum of 75% applied

Maximum		100% power is applied

### Menu item "Boost heater"

No boostheater mounted, inactive

(A secondary heater is used to quickly heat up the bath. This menu item enables or disables the heater.

Standard value:	On)

### Menu item "Time const"

Used to select time to tune. The option precise has to be used to reach maximum temperature accuracy. Options are:

- Fast			60   seconds

- Medium		120 seconds

- Slow 			180 seconds

- Precise		240 seconds

Standard value: 	Precise

### Menu item "Stirrer"

Inactive

(Stirrer 0 .. 100% (step 6%). Inactive)

### Menu item "Low alarm"

min SP to max SP. resolution 0.1°C. No hardware connected, display only

### Menu item "High alarm"

min SP to max SP. resolution 0.1°C. No hardware connected, display only

### Menu item "PID parameter"

PID set 1

PID set 2	-	deactivated

PID set 3	-	deactivated

PID set 4	-	Activated when communication via RS232

Each set offers independant PID settings for:

Proportional band

Integrating

differentiating

Default settings

Pb*	:=	25

I	:=	16

D	:=	0

*Pb= proportional band. Proportional value P is found to be 100/Pb.

### Menu item "Backlight"

On

Off

Standard value:	On

### Menu item "Temp units"

°C

°F

Standard value:	°C

### Menu item "Baudrate"

300

600

1200

2400

4800

9600

19200

38400

Standard value:	9600

### Menu item "Restart"

Restarts system and activates tuning

## PID configuration

The two graphs below show the influence of the integral setting. The graph on top has a setting of 200 for the integral parameter. The bottom graph has a setting of 16. The value of 200 causes less stable control.

I = 200

I = 16

## Manual tuning

The temperature control of the bath is based on a digital PID system. When using different fluids in the bath each with their own heat capacity, the use of external cooling and external connected processes(circulation), or working at different set point temperatures requires new settings of the PID parameters. These parameters have to be optimised after changes to the system when optimal and accurate temperature control of the bath liquid is required.

Tuning of the bath results in:

Stable temperature control of the bath,

No over- or undershoot of the temperature set point,

Quick response to deviations from the set point caused by external disturbances.

- Tuning can be done automatically or manually.

| Parameter | Description | Display |
| --- | --- | --- |
| Proportional band | The bandwidth in display-units over which the output power is proportional between minimum and maximum | P |
| Integration time | Determines the time taken by the controller to remove steady state error signals | I |
| Derivative time | Determines the time taken by the controller to react on error signals. | D |

Table 1 Parameters influencing the temperature control

## Draining bath fluid

Before removing flammable bath liquids take the appropriate fire hazard precautions against these liquids.

Do not remove cold bath liquid. Cold liquid can cause severe burning when spilled. Preheat the bath liquid back to ambient temperature before removing.

When removing bath fluid do not inhale toxic vapor. Always use appropriate ventilation.

### Using the drain tap

The TLC baths can be drained via the drain tap located at the backside of the apparatus.

For safety reasons the tap can only be opened by using a screwdriver. The thread inside the tap is 5/8".

### Fast method:  drain by hand

A.

B.

C.

The bath fluid can easily be removed by using a length of hose. Length of approximately 1.5 mtrs and 10mm inner diameter.

Handle as follows:

Fill the hose with fluid (“A”),

- Close one side of the hose and immerse the other end in the bath (“B”),

- Lower the closed end into a waste bucket and let the bath fluid flow (“C”),

- Remove the last pieces with a sponge or tissue,

- Refill the bath approximately 1 cm below the lid.

## Maintenance

Keep the apparatus free from dust. Regularly check the cooling openings and remove dust with vacuum cleaner. If neccesary remove cover and clean internally. Use appropriate protection when cleaning, dust can be very unhealthy.

# Trouble shooting

## General

| The motor is not running and the lamps are not on. | Check the main voltage and all electrical connections, including switches. Check fuse in wall socket of mains supply. |
| --- | --- |
| Motor is not running, display and heating are OK | Motor overloaded. The motor protection may trip because of mechanical overloading of the pump. Reset fuse by pushing button on rear side of blue motor housing. |
| Motor is not running RED LED blinks on front panel | Overtemperature safety has been activated - Switch off bath - Investigate reason for high bath temperature - Check temperature setting of overtemperature device - Switch bath back on again (wait at least 10 minuse after switching off) |
| Compressor is making “clicking” noise | Bath has been switched on and off too quickly. Wait approximately 10 minutes before switching on again. The clicking noise is a temperature fuse inside the compressor preventing it from overheating. |
| Compressor will not start | When the machine has been switched off the pressure in the cooling circuit can rise to a high level. For this reason the apparatus has to settle the pressure of the cooling liquid for several minutes (5 ... 10) before it can be restarted. When this problem occurs more often it will indicate that dust is blocking the condenser circuit. Remove fron cap and use vacuum cleaner to remove dust. |
| Can not set PID values | Disconnect RS232 connection Select PID parameter set 1 |
| Temperature not stable | Check and set PID parameters. Default is P = 25 I = 16 D = 0 Place top lid on the bath. |
| Temperature does not drop | Bath is not properly insulated. Place top lid on the bath. Insulate circulation circuit Circulation load exceeds heat removal capacity of TLC10. |

# Specifications

## Technical specifications overview

| Item | Unit | TLC10 |
| --- | --- | --- |
| Range | [C] | +60..-10 |
| Setting ± | [°C] | 0,01 |
| Stability* ± | [°C] | 0,05 |
| Uniformity* ± | [°C] | 0,05 |
| Heater | [W] | 1100W |
| Pressure | [mbar] | 270max |
| Capacity | [L/min] | 7 ltr |
| Bath volume | [L] | 3.5 |
| Bath Opening | [mm] | 40x114 |
| Opening effective |  | 32x114 |
| Depth | [mm] | 150 |
| Length | [mm] | 355 |
| Width | [mm] | 195 |
| Height | [mm] | 410 |
| Weight | [kg] | 18 |
| Power (max) | [Watt] | 1250 |
| Voltage | 00T0050 | 230V-50Hz |
|  | 00T0051 | 230V-60Hz |
|  | 00T0052 | 115V-60Hz |

* Absolute min/max value measured over 1hrs in methanol

## TLC10-3 performance

### Stability

Bath fluid is methanol

### Cool down

Bath fluid is methanol

| Heat removal capacity |  |
| --- | --- |
| -10°C | 50W |
| 0°C | 100W |
| 10°C | 150 |
| 20°C | 200 |
| 30°C | 240 |
| 40°C | 280 |
| 50°C | 320 |

# RS232 Communication

## Commands

When TMC70 selected via menu to communicate at Baudrate 4800 use following settings on PC:

4800

8 data bits

1 stop bit

parety none

ST – returns wether bath is stable or not. State 1 = stable, 0 not stable. Limits are set with parameters StableLimitHigh and StableLimitLow defined by ASTM D445. So when PV is in the region of 100°C>PV<0°C StableLimitHigh is defined as 0.05. When the average PV deviation (calculated from a sliding window of 255 seconds) is smaller than StableLimitHigh, the ST command will return stable(1).

When PV is 0°<=PV<=100°C and the average PV (calculated from a sliding window of 255 seconds) is smaller than StableLimitLow, the ST command will return stable(1)

RS - restarts (same as power off-on)

RA – reads raw ADC value, returns 3 bytes binary data, LSB first. This returns the last raw ADC reading. Note that this is updated every 200mS, so reading at a faster rate may return duplicate values. Maximum reading is 0x7FFFFF. This command is used for factory calibration/test.

CT 	<enter> 		Returns the offset value.

CT 	<value> <enter> 	Sets the offset value.

TC – sets/displays time constant in seconds

PL –  Power limit in seconds.

BH – returns boost heater on/off state, 0 or 1

BH<n> - sets boost heater state 0/1

CS – Control state – returns state of control algorithm

bits 4..7 represent the main control stages, bits 0..3 are sub-states within each main state

$00 : control algorithm disabled – allows external PC to take control of  heaters

$10 : startup hold-off started

$11 : startup holdoff in progress

$12 : startup delay finished, decide to boost or cool

$20 : boosting

$30 : Waiting for cool-down

$40 : Start of ratio process

$41 : Ratio process in progress

$42 : waiting for PV<SP-tuneoffset after ratio

$50 : Start of successive approximation process

$51 : Sucessive approximation stabilisation delay

$52 : Successive approximation in progress

$53 : end of sucessuve approximation stage

$60 : Proportional mode   (subject to change - other values to be defined)

CS<n> Sets state of control algorithm

HP : Returns current heater power percentage

HP<nn> : sets heater percentage to <nn> - this command is only useful if control algorithm suspended with CS command.

AC : Returns 1 when a new temperature value is available, 0 if no new has been measured since the last PV or PVH command. Cleared by PV or PVH command. Used for external PC control synchronisation.

PV Returns Proces Value with unit, no decimals. 23C or 74F

PVF returns temp as follows :

Use PVF to return a decimal value with 7 decimal digits, e.g.

12.345342C, works for degrees Centigrade, and degrees Fahrenheid. Only first three are significant. When reading is in degree Fahrenheit character F indicates the unit i.e. 73.3961328F

Whentemperature is negative, minus sign appears i.e. -1.6955468F

PP, returns proportinal band value. Proportional band is 100/gain

PI, returns integrating value

PD, returns derivative value

Setting PID:

Range:

Pb = setting 1..999

I= 0..999

D=0..999

PP followed by value, or space and value, will set proportional band. I.E. PP25 wil set Proportional band to 25.

PI followed by value, or space and value, will set integrating value. I.E. PI16 will set Integratin g value to 16.

PD followed by value , or space and value, will set integrating value. I.E. PD2 will set Derivative value to 2.

The TMC70 contains 4 independent PID sets. Each set can be enabled using the TMC70.EXE program and can have its own name i.e “water”, “silicone oil” or “methanol”. However when the command PP, PI, or PD is used to set a value, automatically set 4 is used. So PID set4 is reserved for RS232 operation and is activated when approached from RS232.

SP	Set point temperature

SP, Set point temperature is returned i.e. 20.22

SP22.50 sets set point temperature to 20.50°C

When controller displays temperature in °F, SP20.50 sets the set point to 20.50°F

SP22.555 results in error code "10"

## RS 232 Cable

D-connector

Seen on “PIN” side (male)

Pin 1

Pin 1

Flatcable 9-pole

D-connector

Seen on BACKSIDE

(female)

# Spare parts

## TLC10-3

| 230 Volt 50HZ | 230V 60HZ | 115 Volt 60Hz | Description |
| --- | --- | --- | --- |
| 25T1295 |  | 25T1301 | Motor for pump |
| 25T1343 |  | n. a. | Capacitor 2uF |
| n. a. |  | 25T1310 | Capacitor 7uF |
| 25T1244 |  | 25T1245 | Fan 120x38mm (Condensor unit) |
| 25T0251 |  | 25T0252 | Heater 1100 Watts |
| 24T8081 |  | n. a. | Motor fuse 0.3 Amp. |
| n. a. |  | 24T8080 | Motor fuse 0, 6 Amp. |
| 04T1001 |  |  | Coupling flexible motor axis |
| 24T8581 |  |  | Mechanical safety thermostat |
| 04T2135 |  |  | Bearing PTFE |
| 23T0012 |  |  | Filling valve |
| 24T8545 |  |  | Mains switch |
| 28T4009 |  |  | Front foil |
| 25T2310 |  |  | PT-100 sensor |
| 06T0512 |  |  | Powerboard PCB (Backpanel) |
| 25T1243 |  |  | Fan 60x60x25 12VDC (Top compartment) |
| 06T0500 |  |  | Microcontroller PCB |
| 06T0502 |  |  | Display PCB |
| 24T8548 |  |  | Plastic Level detector |