January 12, 2011.    Noise characterization of He3 STM.

TEST 1.  Noise measured with spectrum analyzer of RHK.

- Femto DLPCA-200 IV convertor

- DAQ card: PCI-MIO-16XE-10

- RIV convertor = 108 Ohm

- Low noise mode

- Sample is short and grounded at STM

- Temperature sensors OFF

- Input of IV convertor is capped. Output of IV convertor directly connected to RHK

White Noise < 2 fA/sqrt(Hz),  “NULMEETING”

- Same as 1) Now OPAMP connected to current input of RHK. IV convertor grounded at RHK.

Same noise spectrum.

- Same as 1), now TIP connected to input IV convertor. NORMAL OPERATION.

White noise ~ 7 fA/sqrt(Hz), 50 Hz spike of 60 fA/sqrt(Hz), 150 Hz spike of 25 fA/sqrt(Hz),

180-200 Hz spikes up to 50 fA/sqrt(Hz). Height varies in time.

Scope connected to monitor (RHK), does not add noise: on scope 140 mV tt

- Same as 2) Now OPAMP connected at current input of RHK. IV convertor grounded at RHK.

White noise ~ 2 fA/sqrt(Hz), 50 Hz spike of 55 fA/sqrt(Hz), 150 Hz of 30 fA/sqrt(Hz),

180-200 Hz spikes up to 20 fA/sqrt(Hz). Height varies in time.

Scope connected to monitor (RHK), does not add noise: on scope 60 mV tt.

OPAMP STRONGLY REDUCES WHITE NOISE! ALSO, a 2 kHz low pass filter gives similar reduction.

- Same as 4), now temperature controllers are OFF.

White noise ~ 3 fA/sqrt(Hz), 50 Hz spike of 25 fA/sqrt(Hz), 150 Hz weg,

180-200 Hz spikes up to 25 fA/sqrt(Hz). Height varies in time.

- Same as 3) Now IV convertor grounded at STM. Similar noise spectrum as 3)

Scope connected to monitor (RHK), does not add noise: on scope 60 mV tt

- Now OPAMP connected to output IV convertor

Sample is grounded at STM.

No difference in noise compared to 4)

Found Origin of 50 Hz Noise.

- Temperature sensors couple in, especially lower ITC controller.

Putting off decreases 50 Hz from 600 to 140 fA/sqrt(Hz) (no opamp)

20 fA/sqrt(Hz) (with OPAMP)

- Grounding Sample wire to STM gets rids of 50 Hz. Now 200 Hz spike still present, at 160 fA/sqrt(Hz)

- Grounding Sample wire to RHK gets rids of 50 Hz. Now 200 Hz spike still present, at 160 fA/sqrt(Hz)

NOISE LEVEL WHEN

- TIP is grounded at STM

- OPAMP is grounded at RHK

- Temperature sensors OFF and grounded at STM.

- 50 Hz: 17 fA/sqrt(Hz)

- 100 Hz: 15 fA/sqrt(Hz)

- 140 Hz: 10 fA/sqrt(Hz)

- 180 Hz: 12 fA/sqrt(Hz)

- White noise < 3 fA/sqrt(Hz)

Basically, all spikes are below 17 fA/sqrt(Hz)

TEST 2

Test 2:

- Femto DLPCA-200 IV convertor

- DAQ card: PCI-MIO-16XE-10

- RIV convertor = 107 Ohm

- Low noise mode

- IV convertor at STM, 5 meter cable to DAQ card. No current input (capped).

IV convertor is grounded at the STM.

Output of IV convertor has instrumentation amplifier just before DAQ to disconnect shield.

Measurement: rate of 1 kHz, 1 kSamples.

Noise <1 mV (4 bits)

- Same as 1), now without instrumentation amp.

Noise ~ 2 mV

- Same as 1), now with TIP connected to current input of STM.

Connected to SAMPLE gives same result.

Note: Piezo’s are grounded.

Noise < 7 mV  (which is <1 nA).

- Same as 3), now also connected to RHK.

Noise  < 7 mV

- As 3), now only grounded at DAQ

Noise ~ 7 mV

- Now connected to DAQ without instrumentation amplifier.

(grounded at STM & RHK)

Noise ~ 20 mV

- Connected to DAQ and RHK without instrumentation amplifier.

(grounded at STM & RHK)

Noise ~ 30 mV

- Start measuring the monitor output of RHK. Directly connected to DAQ.

Current input of RHK is output of IV convertor

(IV convertor is grounded at RHK box)

Noise ~ 12 mV

- Same as 8), now operational amplifier connected to RHK and IV convertor grounded just before opamp.

OPAMP works well!

Noise ~ 7 mV.

- Same as 9), now IV convertor grounded at STM.

Seems a bit improved:

Noise ~ 6 mV.

BEST OPTION!!!